Analysis of Neutron-Induced Degradation and Recovery Characteristics of

LaBr3(Ce) Detector Performance in PGNAA Environments

Ki-Yoon Lee1,2, Kyung-Hwan Jung1, Hyun-Dong Kim1, Hyun-Soo Ha1, Jung-Hyeon Eo2, Jun-Young Shin2, Jae-Hyeon Kim3, Kyung-Min Oh3, Jin-Hyung Park3,*, Cheol-Ha Baek1,4,*

1Department of Health Medical Science, Kangwon National University, Samcheok, Gangwon-do, Republic of Korea

23I Solution Inc., Suwon, Gyeonggi-do, Republic of Korea

3Advanced Radiation Technology Institute, Korea Atomic Energy Research Institute, Jeongeup, Jeonbuk-do, Republic of Korea

4Department of Radiological Science, Kangwon National University, Samcheok, Gangwon-do, Republic of Korea

Corresponding Author Email: jhpak@kaeri.re.kr, baekch@kangwon.ac.kr

Figure 1: Comparison of silicon (Si) prompt gamma-ray spectra in activated (red) and recovered (blue) states of the LaBr₃(Ce) detector following neutron irradiation.

LaBr₃(Ce) detectors are widely used in gamma-ray spectroscopy owing to their superior energy resolution and high light yield. However, in high neutron-flux environments, intrinsic activation of the detector constituents results in a rapid increase in background radiation and degradation of spectral integrity, thereby impeding precise spectroscopic analysis. This study quantitatively evaluates the degradation and recovery kinetics of LaBr₃(Ce) performance following neutron irradiation and assesses its subsequent impact on the detection of the 1778.2 keV silicon (Si) prompt gamma-ray peak[1].

Detector performance was monitored over a 25-day post-irradiation period using a ⁶⁰Co check source and intrinsic background analysis. During the initial "Activated Phase" (Days 3 to 18), the energy resolution at 1332.5 keV deteriorated from 3.4% to over 21.0%, while the background index (BgIdx) exhibited an approximately six-fold increase compared to the recovery state. In this period, the silicon peak at 1778.2 keV remained undetectable, as the signal-to-background ratio (SBR) approached 1 due to dominant activation-induced noise.

In contrast to the earlier phase, a distinct recovery of spectroscopic performance was observed as the cooling time progressed. Starting from Day 23, detector performance began to restore; by Day 25, the intrinsic background similarity reached 1.00 relative to the pre-irradiation state, and the energy resolution at 1461 keV recovered to 2.09%. During this recovery phase, precise detection and analysis of the silicon prompt gamma-ray peak, which had been previously obscured, became feasible again. These results delineate the operational constraints of LaBr₃(Ce) detectors in high-neutron environments and propose an optimal cooling time for reliable prompt gamma neutron activation analysis (PGNAA) in silicon-based applications.

1. Demidov, A. M., Govor, L. I., Cherepantsev, Yu. K., Ahmed, M. R., Al-Najjar, S., Al-Amili, M. A., Al-Assafi, N., Rammo, N., Atlas of gamma-ray spectra from the inelastic scattering of reactor fast neutrons. Moscow: Atomizdat, (1978) pp. 77~81.

This research was supported by the Technology Development Program of the Ministry of SMEs and Startups [RS-2024-00469068].